## Slide 1: Status of the S-DALINAC
Gefördert durch die DFG im Rahmen des SFB 1245 und des GRK 2128
Picture: Jan-Christoph Hartung

## Slide 2
S-DALINAC
Superconducting-DArmstadt-LINear-ACcelerator
Beam current: up to 20 µA
Beam energy: up to 130 MeV
Thrice recirculating operation
Energy gain injector:
Energy gain LINAC:
7.6 MeV
30.4 MeV
Thermionic Gun
5 m
Spin-Polarized Electron Gun
Injector
LINAC
DHIPS
Recirculation Beam Lines
NEPTUN
QCLAM Magnet Spectrometer
169° LINTOTT Magnet Spectrometer
180° Setup@ QCLAM

## Slide 3: Milestones Since August 2017
A1SC01
ERL-Cup
E0F1-Cup
Thermionic Gun
Injector LINAC
Main LINAC
Recirculations
First ERL in Germany!
Energy gain injector:
Energy gain LINAC:
Current (to injector):
2.5 MeV
20.0 MeV
~ 1.2 µA
Physik Journal 10/2017; publication in preparation
August 17

## Slide 4: Milestones Since August 2017
169° LINTOTT Magnet Spectrometer
High-Energy Scraper-System
First commissioning of high-energy scraper-system: Test of energy-scraper and horizontal halo-scraper
Energy-scraper
Halo-scraper
vertical
Halo-scraper horizontal
+ vertical
7.5 m
Lars Jürgensen
August 17
October / November 17

## Slide 5: Optimization of Energy Resolution
Linewidth of elastic peak in 169°-spectrometer:
42.5 MeV on 1 mg/cm² gold foil at 93°
Duration of each measurement approx. 5 min
Lars Jürgensen
  Fluctuation of beam energy
Notes: Gemessen wurde die Breite der elastischen Linien, wobei elektronen mit einer Strahlenergie von 42,5 MeV auf ein Goldtarget mit einer Massenbelegung von 1 mg/cm² geschossen wurden. Jeder Messpunkt entspricht einer 5 minütigen Datenaufnahme. Die Grafik zeigt in schwaz die Linienbreite in keV und in Rot und blau zwei verschiedene Methoden zur Bestimmung der Transmission. Die Transmission und die Linienbreite nehmen beim Schließen des Energiescrapers erwartungsgemäß ab. Bei noch großen Spaltbreiten führen Schwankungen der Strahlenergie die in die 5 minütige Messdauer fallen zu starken abweichungen der emittelten Linienbreite.

## Slide 6: Best Result – New Record
Combination of energy- and horizontal halo-scraper:
Linewidth without scraper: 25 keV
Linewidth with scraper less than 10 keV!  7.9 keV at 42.5 MeV
Lars Jürgensen
(e,e') with energy resolution of gamma spectroscopy (3 – 10 keV)
Notes: Bei einzelnen Versuchen zur Kombination von Energiescraper und Haloscraper konnten Linienbreiten unter 10 keV gemessen werden.
Die kleinste erreichte Linienbreite betrug 7,9 keV bei einer Strahlenergie von 42,5 MeV.
Damit bewegen wir uns in der Größenordnung der Detektorauflösung!
Die zugehörige Messung ist hier im Vergleich zu einer Messung ohne Scrapereinsatz gezeigt. Aufgetragen ist die normierte Häufigkeit über dem Energieverlust der gestreuten Elektronen.
Bei einer weiteren Verbesserung des Elektronenstrahls konvergiert die Linienbreite zwangsläufig gegen die verbleibenden Beiträge von Spektrometer und Target.
Detektorauflösung (4 keV) Spektrometermagnet (6,5 keV) und Targetverbreiterung (1 keV)
5 minuten gegen 20 minuten messzeit
25 keV gegen 8 keV
1-2 % Transmission
Messung bei 1uA, System ausgelegt für 20uA , -> Vervierfachung des Stromes bringt messergebnis auch in 5 minuten

## Slide 7: Reduction of Background
Only energy-scraper
Lars Jürgensen
Notes: Man kann sowohl bei dem Goldtarget als auch beim Leertarget eine deutliche Reduktion des Photonenuntergrunds sehen.

## Slide 8: Milestones Since August 2017
169° LINTOTT Magnet Spectrometer
  Determination of beam size
18 cm
To spectrometer
To camera
Beam
Gold target
OTR-target
Wire scanner
Manuel Dutine
August 17
October / November 17

## Slide 9: Beam Width - Wire Scanner
6
5
4
3
2
1
Count rate in events/s
0.5
1
2
1.5
2.5
Horizontal wire
45°-wire
Position in mm
Position in mm
11
12
13
14
15
44
45
46
47
48
0
0
Manuel Dutine
Vertical beam width in mm (measured) | Beam width in mm (measured) | Horizontal beam width in mm (calculated)

## Slide 10: Beam Width – OTR Target
1 mm
1 mm
Vertical beam width: mm
Horizontal beam width: ) mm
Statistical uncertainties one order of magnitude smaller than wire scanner:
  ) mm
  ) mm
Manuel Dutine

## Slide 11: Measurement of Dispersion
Dispersion: (
Large uncertainties due to cross-fade and brigth scale
Future: Higher precision with OTR-measurement-system
5 mm
Manuel Dutine
BeO Target

## Slide 12: Milestones Since August 2017
  Very good setting
  24.11.2017: ~98 % transmission to extraction at design current for recirculating operation (20 µA)  first reach of design value
  Beam losses very small
BeO target
10 mm
August 17
October / November 17

## Slide 13: Milestones Since August 2017
Thermionic Gun
5 m
Injector
  Beam based alignment injector linac
F. Hug et al., IPAC 2018, in press
  5-cell cavity
  BeO target
August 17
October / November 17
December 17

## Slide 14: Beam Based Alignment
 |  | εnx,rms [π µm] | εny,rms [π µm]
Before alignment | simulated | 0.54 | 0.31
 | measured | 1.3 ± 0.1 | 0.4± 0.02
After alignment | simulated | 0.43 | 0.24
 | measured | 0.84± 0.05 | 0.15± 0.05
  Before alignment
  After alignment
  Much less steering effect of 5-cell cavity
  Better focus at optimum rf phase
  Total change of rf phase of more than 20° possible without losing beam on screen
F. Hug et al., IPAC 2018, in press

## Slide 15: Milestones Since August 2017
August 17
October / November 17
December 17
Upgrade path length adjustment systems of first and third recirculation
Work on vacuumsystem (e.g. pumps, new elements)
Repair of power supplies, rf-amplifiers
Interlock system (personell and machine safety): small corrections (not relevant to safety), upgrades to increase user-friendliness
Alignment: cooperation contract with Franfurt UAS in preparation
…
Refurbishment of He-refrigerator, work on SRF-modules
Shut Down

## Slide 16: Simplified Scheme
Kaeser compressor (in operation approx. 9-11.5 bar)
Cryostat
Buffer
(50 m³, 10 bar)
Sauer compressor
(up to 200 bar)
Bypass-
valve
„Lade“-valve
„Ablade“-valve
Low pressure
High pressure
Balloon
HP-bottles
(2.55 m³, 200 bar)
Plant
HX 1
HX 2
HX 3
Residual gas
analyzer
HX 4
HX 6
HX 5
HX 7
HX 8
HX 9
Turbine 2
Turbine 1
Leybold pumping stages
(35 mbar to 1050 mbar)
Machine circuit
Dirty circuit
Exchange buffer

## Slide 17: Simplified Scheme
Kaeser compressor (in operation approx. 9-11.5 bar)
Cryostat
Buffer
(50 m³, 10 bar)
Sauer compressor
(up to 200 bar)
Bypass-
valve
„Lade“-valve
„Ablade“-valve
Low pressure
High pressure
Balloon
HP-bottles
(2.55 m³, 200 bar)
Plant
HX 1
HX 2
HX 3
Residual gas
analyzer
HX 4
HX 6
HX 5
HX 7
HX 8
HX 9
Turbine 2
Turbine 1
Leybold pumping stages
(35 mbar to 1050 mbar)
Machine circuit
Dirty circuit
Exchange buffer
Kaeser
(Year of construction: 2003)
Leybold
(Year of construction: 2004)
Bauer
(Year of construction: ~1984)
Sulzer cryo plant
(Year of construction: 1982)
Linde

## Slide 18: Costs for Refurbishment
Turbines, residual gas analyzer
RAK, spare parts cryo plant
Kaeser compressor, motor fan
Leybold pumping stages
Sauer compressor
Total
27.0 k€
64.2 k€
54.1 k€
70.9 k€
33.5 k€
249.7 k€
Investment of IKP to prepare for an operation of the cryo plant in the next years:
Experience up to now: No internal / unwanted bypass gas flows any more  high power of system is visible

## Slide 19: Status He-Refrigerator
No problems in cooldown after complete disassembly and reassembly of nearly every element! All degraded parts replaced.
Comparison of cooling parameters before (operational setting) and after refurbishment (cooldown setting)
 | HP in bar | Position Bypass in % | Cooling Power in W | Differential pressure increase in bar per week
Before Refurbishment | 11.5 | ~ 15 | ~ 80 | ~ 0
After Refurbishment | 10.0 | ~ 53 | ~ 55 | ~ 0.04
Important: Refrigerator will be performance-checked and further pushed for cooling power in the next days  no problems for beam time expected

## Slide 20: Status Cavities – Measurement ongoing Average Q0 vs max. Gradient
A1SC01 and A1SC02: field flatness tuning @ IKP,
ultrasonic cleaning with ultrapure water @ DESY)
Uncertainties in range ~1% to ~40% due to measurement
Ruben Grewe

## Slide 21
Beam Time Schedule 2018
70 MeV non-isochronous
Beam tuning, machine experiments, KW 37 installation of He-target
Neptun Tagger
Q-CLAM
Q-CLAM He-Target
Cooldown, tests cavities and He-refrigerator
Q-CLAM 180°
Diamond detector tests
Q-CLAM 180° (35 MeV single pass)
Twice recirculating ERL

## Slide 22: Thank you for your attention!
Picture: Jan-Christoph Hartung